## Load Operation Description
UART Command | Description | Load Response
&ModeI; | Configure the electronic load to constant current mode (defaul mode) | Constant Current mode is active. | Load Maximum and minimum allowed electrical values.
&ModeR; | Configure the electronic load to constant resistance mode | Constant Resistance mode is active. | Variable | Minimum Value | Maximum Value
&ModeP; | Configure the electronic load to constant power mode | Constant Power mode is active. | Load Current | 0 | 30 A
&current xx.xx; | Set the reference value (xx.xx) for the constant current mode | No response | Load Resistance | 0,1 Ohms | 10000 Ohms
&resistor xx.xx; | Set the reference value (xx.xx) for the constant resistance mode | No response | Load Power | 0 W | 320 W
&power xx.xx; | Set the reference value (xx.xx) for the constant power mode | No response | Load Voltage | 0 V | 50 V
&FanEN; | The cooling fan of the load is enabled | Fan is enabled. | Note: in case of a user fault everything is also limited in the firmware
&FanDIS; | The operation of the fan is controlled by the reading values of the thermocouples. | Fan is automatically controlled.
&GenFault; | Generate fault condition. All circuits are off and the cooling fan is ON | Fault in manually generated.
&CLRFault; | Clear fault condition | Fault is cleared. | Electronic Load Curve
User LED function | Description
Blinking period 2 seconds | Normal operation, while all the temperature readings are below 45 °C. The cooling fan is off.
Blinking period 0.5 seconds | Normal operation, while at least one of the temperature readings is over 70 °C. The cooling fan is ON.
Constanly ON | The load is deactivated because of the overtemperature protection at 110 °C. The condition is latched and the fan is constantly ON until a CLRFault command is received.
Every 1 second the electronic load sends via the UART bus a monitoring message. The format of the message is shown below
V_Load = xx.xx, I_Load = xx.xx, P_Load = xx.xx, Max_Temp = xx.xx, Mode_XX, Ref_xx.xx
V_Load | The voltage reading across the electronic load terminals
I_Load | The total current consuption of the electronic load
P_Load | The power consumption of the load
Max_Temp | The maximum temperature reading of the load
Mode_XX | The mode of operation of the load (CR for constant resistance, CP for constant power, CC for constant current)
Ref | The reference command value for the coresponding mode of operation

## Power_Supplies
Main 5V power supply
DC coupling | V = 5.13V
The BUCK converter operates under ligh load, and thus DCM is the mode of operation or diode emulation mode with reduced frequency.
AC Coupling | 64 mV ptp
3.3 V rail |  |  | 3 V rail
DC coupling | 3.26 V |  | DC coupling |  | 3 V
AC coupling |  |  | AC coupling
Control op amp for current source Vcc |  | RC filter for Vcc filtering with Fc = 72 kHz
AC coupling | 45mV ptp

## PowerCircuit Steady State
Load reference 6 A (1 A per circuit)
Current measured with the shunt resistor (circuit 4)
DC coupling | AC coupling
Load reference 12 A (2 A per circuit)
Current measured with the shunt resistor (circuit 4)
DC coupling | AC coupling
Load reference 18 A (3 A per circuit)
Current measured with the shunt resistor (circuit 4)
DC coupling | AC coupling
Load reference 23 A (3,83 A per circuit)
Current measured with the shunt resistor (circuit 4)
DC coupling | AC coupling

## PowerCircuit Transients
Dynamic response
Load current is controlled by utilizing both analog and digital closed loop control.
CC Mode | CR Mode | CP Mode
Reference command from 6A to 23 A (load voltage constant at 12 V) | Reference command from 2 Ohms to 1.5 Ohms | Reference command from 100 W to 200 W
yellow circuit current | yellow load voltage | yellow load voltage
| cyan circuit current | cyan circuit current
Reference command from 23 A to 6 A (load voltage constant at 12 V) | Reference command from 1.5 Ohms to 2 Ohms | Reference command from 200 W to 100 W
yellow circuit current | yellow load voltage | yellow load voltage
| cyan circuit current | cyan circuit current
Constant 18 A reference with step changes to the load voltage | Constant 1.8 Ohm reference with step changes to the load voltage | Constant 140 W reference with step changes to the load voltage
yellow load voltage | yellow load voltage | yellow load voltage
cyan circuit current | cyan circuit current | cyan circuit current
purple circuit power

## FanControl
If the temperature is below 45 °C then the cooling fan is OFF. | Fans Used
If the temperature is above 70 °C the colling fan is ON. | 2x MGT3612HB-W28 (12 V/ 0.2 A)
The cooling fan can also be enabled by sending the &FanEN; command via UART.
The cooling fan operated automatically according to the temperature, as described above, by sending the &FanDIS; command.
In the case of the overtemperature fault, the fan is enabled and remains at the ON state until the CLR fault command. After the CLRFault command, the fan is automatically controlled according to the temperature readings.
Every time that the fan is enabled the duty cycle ramps up from 0 to 100 % in 1 second as shown below.
Fan ramp up time

## Temperature measurement
Temperature | Resistance (kOhms) | RpEq(kOhms) | Vdiv_Out(V) | P_NTC (mW) | Power OK | ADC Value | Calc temp | Dif | Vin (V) | 3.3 | NTC part number | NRL1104J3435B2F
-40 | 200.79 | 66.7542 | 1.1923 | 7.0803 | check | 1627.5372 | -24.7279 | 15.2721 | R1 (kOhms) | 118 |  | crimped on 3 mm round terminal and potted with Arctic MX-4 tthermal paste
-35 | 152.4 | 60.3803 | 1.117 | 8.1873 | check | 1524.7379 | -24.2312 | 10.7688 | Rp (kOhms) | 100
-30 | 116.74 | 53.8618 | 1.0342 | 9.1624 | check | 1411.7179 | -22.9533 | 7.0467 | Dissipation factor NTC (mW/C) | 1.8
-25 | 90.07 | 47.3878 | 0.94553 | 9.926 | check | 1290.6536 | -20.7344 | 4.2656
-20 | 70.09 | 41.2076 | 0.85414 | 10.4088 | check | 1165.8967 | -17.5277 | 2.4723
-15 | 55.02 | 35.4922 | 0.76306 | 10.5828 | check | 1041.5812 | -13.4034 | 1.5966
-10 | 43.5 | 30.3136 | 0.67448 | 10.4581 | check | 920.6679 | -8.5025 | 1.4975
-5 | 34.63 | 25.7224 | 0.59061 | 10.0728 | check | 806.1817 | -3.0536 | 1.9464
0 | 27.74 | 21.716 | 0.51292 | 9.4839 | check | 700.1322 | 2.6954 | 2.6954
5 | 22.32 | 18.2472 | 0.44196 | 8.7513 | check | 603.2755 | 8.5357 | 3.5357
10 | 18.09 | 15.3188 | 0.37918 | 7.948 | check | 517.5837 | 14.172 | 4.172
15 | 14.76 | 12.8616 | 0.32434 | 7.127 | check | 442.721 | 19.4566 | 4.4566
20 | 12.11 | 10.8019 | 0.27675 | 6.3247 | check | 377.7671 | 24.3142 | 4.3142
25 | 10 | 9.0909 | 0.23605 | 5.572 | check | 322.2103 | 28.6699 | 3.6699
30 | 8.29 | 7.6554 | 0.20105 | 4.8758 | ok | 274.4301 | 32.564 | 2.564
35 | 6.92 | 6.4721 | 0.17159 | 4.2547 | ok | 234.2187 | 35.9474 | 0.94741
40 | 5.81 | 5.491 | 0.14673 | 3.7058 | ok | 200.2907 | 38.8776 | -1.1224
45 | 4.898 | 4.6693 | 0.12561 | 3.2214 | ok | 171.4598 | 41.4219 | -3.5781
50 | 4.1544 | 3.9887 | 0.1079 | 2.8025 | ok | 147.2847 | 43.5937 | -6.4063
55 | 3.5333 | 3.4127 | 0.092758 | 2.4351 | ok | 126.6143 | 57.4091 | 2.4091
60 | 3.0173 | 2.9289 | 0.079927 | 2.1172 | ok | 109.1 | 58.5198 | -1.4802
65 | 2.5869 | 2.5217 | 0.069046 | 1.8429 | ok | 94.2474 | 62.3461 | -2.6539
70 | 2.2261 | 2.1776 | 0.059796 | 1.6062 | ok | 81.6217 | 67.6802 | -2.3198
75 | 1.9225 | 1.8862 | 0.051921 | 1.4022 | ok | 70.8718 | 73.7296 | -1.2704
80 | 1.6661 | 1.6388 | 0.045203 | 1.2264 | ok | 61.702 | 79.9858 | -0.014247
85 | 1.4485 | 1.4278 | 0.039453 | 1.0746 | ok | 53.8535 | 86.1419 | 1.1419
90 | 1.2651 | 1.2493 | 0.034572 | 0.94476 | ok | 47.1906 | 91.9481 | 1.9481
95 | 1.1085 | 1.0963 | 0.030378 | 0.83251 | ok | 41.4664 | 97.3619 | 2.3619
100 | 0.9743 | 0.9649 | 0.026766 | 0.73529 | ok | 36.535 | 102.341 | 2.341
105 | 0.8588 | 0.85149 | 0.023642 | 0.65085 | ok | 32.2716 | 106.881 | 1.881
110 | 0.7594 | 0.75368 | 0.020944 | 0.57761 | ok | 28.5881 | 110.9791 | 0.9791
115 | 0.6735 | 0.66899 | 0.018604 | 0.51388 | ok | 25.394 | 114.6644 | -0.33563
120 | 0.5988 | 0.59524 | 0.016563 | 0.45813 | ok | 22.6083 | 117.9785 | -2.0215
125 | 0.5338 | 0.53097 | 0.014783 | 0.40937 | ok | 20.1781 | 120.9457 | -4.0543
Points of temperature probing

## Load operation Data logging
Time (s) | Vload (V) | I_Load(A) | PLoad(W) | Max_Temperature (°C) | Resistance
1 | 11.99 | 0.03 | 0.42 | 36.9 | 399.6667
2 | 11.99 | 0.03 | 0.41 | 36.95 | 399.6667
3 | 11.99 | 0.03 | 0.43 | 36.86 | 399.6667
4 | 11.99 | 0.03 | 0.39 | 36.83 | 399.6667
5 | 11.99 | 0.02 | 0.32 | 36.95 | 599.5
6 | 11.99 | 0.03 | 0.39 | 36.9 | 399.6667
7 | 11.99 | 0.03 | 0.39 | 36.87 | 399.6667
8 | 11.99 | 0.02 | 0.35 | 36.87 | 599.5 | A/A | Modes and references
9 | 11.99 | 0.03 | 0.4 | 36.71 | 399.6667 | 1 | Constant current mode
10 | 11.99 | 0.03 | 0.41 | 36.83 | 399.6667 | 2 | Constant current mode with 10 A reference
11 | 11.99 | 0.03 | 0.37 | 36.8 | 399.6667 | 3 | Constant current mode with 15 A reference
12 | 11.99 | 0.03 | 0.38 | 36.79 | 399.6667 | 4 | Constant resistance mode
13 | 11.99 | 0.03 | 0.42 | 36.76 | 399.6667 | 5 | Resistor is 2 Ohms
14 | 11.99 | 0.03 | 0.38 | 36.73 | 399.6667 | 6 | Resistor is 1.5 Ohms
15 | 11.99 | 0.02 | 0.34 | 36.71 | 599.5 | 7 | Constant power mode
16 | 11.52 | 10.12 | 116.68 | 36.58 | 1.1383 | 8 | Power is 50 W
17 | 11.52 | 10.12 | 116.68 | 36.55 | 1.1383 | 9 | Power is 100 W
18 | 11.51 | 10.12 | 116.67 | 36.57 | 1.1374 | 10 | Constant current mode
19 | 11.52 | 10.12 | 116.67 | 36.54 | 1.1383 | 11 | Constant current mode with 15 A reference
20 | 11.51 | 10.12 | 116.64 | 36.58 | 1.1374 | 12 | OVT fault
21 | 11.51 | 10.12 | 116.63 | 36.63 | 1.1374 | 13 | Fault cleared
22 | 11.52 | 10.12 | 116.65 | 36.75 | 1.1383
23 | 11.52 | 10.12 | 116.73 | 36.85 | 1.1383
24 | 11.52 | 10.12 | 116.68 | 37.01 | 1.1383
25 | 11.52 | 10.12 | 116.7 | 37.15 | 1.1383
26 | 11.51 | 10.12 | 116.66 | 37.33 | 1.1374
27 | 11.52 | 10.12 | 116.67 | 37.52 | 1.1383
28 | 11.52 | 10.12 | 116.72 | 37.71 | 1.1383
29 | 11.52 | 10.12 | 116.69 | 37.92 | 1.1383
30 | 11.53 | 10.12 | 116.76 | 38.15 | 1.1393
31 | 11.52 | 10.12 | 116.74 | 38.36 | 1.1383
32 | 11.52 | 10.12 | 116.67 | 38.59 | 1.1383
33 | 11.52 | 10.12 | 116.73 | 38.81 | 1.1383
34 | 11.52 | 10.12 | 116.71 | 39.05 | 1.1383
35 | 11.52 | 10.12 | 116.67 | 39.3 | 1.1383
36 | 11.52 | 10.12 | 116.73 | 39.51 | 1.1383
37 | 11.51 | 10.12 | 116.6 | 39.78 | 1.1374
38 | 11.52 | 10.12 | 116.7 | 40.03 | 1.1383
39 | 11.52 | 10.12 | 116.71 | 40.22 | 1.1383
40 | 11.52 | 10.11 | 116.63 | 40.49 | 1.1395
41 | 11.52 | 10.12 | 116.71 | 40.71 | 1.1383
42 | 11.51 | 10.12 | 116.6 | 40.94 | 1.1374
43 | 11.52 | 10.13 | 116.75 | 41.16 | 1.1372
44 | 11.51 | 10.12 | 116.61 | 41.35 | 1.1374 | A/A | Modes and references
45 | 11.51 | 10.12 | 116.57 | 41.57 | 1.1374 | 1 | Constant current mode
46 | 11.51 | 10.13 | 116.7 | 41.79 | 1.1362 | 2 | Constant current mode with 10 A reference
47 | 11.51 | 10.13 | 116.74 | 42.01 | 1.1362 | 3 | Constant current mode with 15 A reference
48 | 11.51 | 10.11 | 116.49 | 42.22 | 1.1385 | 4 | Constant resistance mode
49 | 11.51 | 10.11 | 116.49 | 42.47 | 1.1385 | 5 | Resistor is 2 Ohms
50 | 11.51 | 10.12 | 116.64 | 42.64 | 1.1374 | 6 | Resistor is 1.5 Ohms
51 | 11.51 | 10.12 | 116.62 | 42.82 | 1.1374 | 7 | Constant power mode
52 | 11.51 | 10.12 | 116.58 | 42.99 | 1.1374 | 8 | Power is 50 W
53 | 13.5 | 10.12 | 136.8 | 43.17 | 1.334 | 9 | Power is 100 W
54 | 17.51 | 10.12 | 177.41 | 43.35 | 1.7302 | 10 | Constant current mode
55 | 17.51 | 10.12 | 177.35 | 43.53 | 1.7302 | 11 | Constant current mode with 15 A reference
56 | 17.51 | 10.12 | 177.33 | 43.7 | 1.7302 | 12 | OVT fault
57 | 17.5 | 10.12 | 177.29 | 43.89 | 1.7292 | 13 | Fault cleared
58 | 17.51 | 10.12 | 177.3 | 44.09 | 1.7302
59 | 18.5 | 10.12 | 187.46 | 44.28 | 1.8281
60 | 19.5 | 10.12 | 197.52 | 44.45 | 1.9269
61 | 19.51 | 10.12 | 197.56 | 44.68 | 1.9279
62 | 19.5 | 10.12 | 197.55 | 44.86 | 1.9269
63 | 19.5 | 10.12 | 197.48 | 45.06 | 1.9269
64 | 19.5 | 10.12 | 197.55 | 45.26 | 1.9269
65 | 19.51 | 10.12 | 197.59 | 45.49 | 1.9279
66 | 19.5 | 10.12 | 197.49 | 45.69 | 1.9269
67 | 19.5 | 10.12 | 197.57 | 45.9 | 1.9269
68 | 19.51 | 10.12 | 197.54 | 46.13 | 1.9279
69 | 19.5 | 10.12 | 197.55 | 46.34 | 1.9269
70 | 18.51 | 10.12 | 187.49 | 46.52 | 1.8291
71 | 17.51 | 10.12 | 177.31 | 46.73 | 1.7302
72 | 16.51 | 10.12 | 167.23 | 46.93 | 1.6314
73 | 15.51 | 10.12 | 157.05 | 47.16 | 1.5326
74 | 14.51 | 10.12 | 146.94 | 47.33 | 1.4338
75 | 13.51 | 10.12 | 136.86 | 47.54 | 1.335 | A/A | Modes and references
76 | 13.5 | 10.12 | 136.71 | 47.71 | 1.334 | 1 | Constant current mode
77 | 11.59 | 10.12 | 117.38 | 47.89 | 1.1453 | 2 | Constant current mode with 10 A reference
78 | 10.26 | 10.12 | 103.91 | 48.07 | 1.0138 | 3 | Constant current mode with 15 A reference
79 | 5.59 | 10.12 | 56.65 | 58.54 | 0.55237 | 4 | Constant resistance mode
80 | 3.5 | 10.12 | 35.53 | 61.33 | 0.34585 | 5 | Resistor is 2 Ohms
81 | 2.51 | 10.12 | 25.45 | 61.83 | 0.24802 | 6 | Resistor is 1.5 Ohms
82 | 2.51 | 10.12 | 25.46 | 62.24 | 0.24802 | 7 | Constant power mode
83 | 2.51 | 10.12 | 25.46 | 62.65 | 0.24802 | 8 | Power is 50 W
84 | 3.4 | 10.12 | 34.43 | 63.07 | 0.33597 | 9 | Power is 100 W
85 | 5.5 | 10.12 | 55.71 | 63.34 | 0.54348 | 10 | Constant current mode
86 | 7.42 | 10.12 | 75.18 | 63.64 | 0.7332 | 11 | Constant current mode with 15 A reference
87 | 10.51 | 10.12 | 106.41 | 63.93 | 1.0385 | 12 | OVT fault
88 | 11.5 | 10.12 | 116.51 | 64.06 | 1.1364 | 13 | Fault cleared
89 | 11.51 | 10.12 | 116.57 | 64.35 | 1.1374
90 | 11.51 | 10.12 | 116.54 | 64.4 | 1.1374
91 | 11.51 | 10.12 | 116.6 | 64.65 | 1.1374
92 | 11.5 | 10.12 | 116.48 | 64.86 | 1.1364
93 | 11.51 | 10.12 | 116.57 | 64.88 | 1.1374
94 | 11.51 | 10.12 | 116.56 | 65.28 | 1.1374
95 | 11.28 | 15.18 | 171.37 | 65.22 | 0.74308
96 | 11.29 | 15.19 | 171.64 | 65.41 | 0.74325
97 | 11.29 | 15.18 | 171.57 | 65.69 | 0.74374
98 | 11.29 | 15.18 | 171.61 | 65.81 | 0.74374
99 | 11.3 | 15.18 | 171.72 | 66.03 | 0.7444
100 | 11.3 | 15.18 | 171.65 | 66.42 | 0.7444
101 | 11.3 | 15.19 | 171.74 | 66.88 | 0.74391
102 | 11.3 | 15.18 | 171.68 | 67.1 | 0.7444
103 | 11.3 | 15.18 | 171.69 | 67.43 | 0.7444
104 | 11.3 | 15.19 | 171.74 | 68 | 0.74391
105 | 11.3 | 15.18 | 171.65 | 68.46 | 0.7444
106 | 11.3 | 15.19 | 171.73 | 68.91 | 0.74391
107 | 11.3 | 15.18 | 171.7 | 69.56 | 0.7444 | A/A | Modes and references
108 | 11.3 | 15.18 | 171.67 | 69.93 | 0.7444 | 1 | Constant current mode
109 | 11.3 | 15.19 | 171.74 | 70.17 | 0.74391 | 2 | Constant current mode with 10 A reference
110 | 11.3 | 15.18 | 171.67 | 69.45 | 0.7444 | 3 | Constant current mode with 15 A reference
111 | 11.31 | 15.18 | 171.82 | 69.64 | 0.74506 | 4 | Constant resistance mode
112 | 11.31 | 15.18 | 171.82 | 70.23 | 0.74506 | 5 | Resistor is 2 Ohms
113 | 11.31 | 15.18 | 171.87 | 70.67 | 0.74506 | 6 | Resistor is 1.5 Ohms
114 | 11.31 | 15.19 | 171.95 | 71.28 | 0.74457 | 7 | Constant power mode
115 | 11.31 | 15.18 | 171.77 | 71.81 | 0.74506 | 8 | Power is 50 W
116 | 11.31 | 15.18 | 171.9 | 72.68 | 0.74506 | 9 | Power is 100 W
117 | 11.31 | 15.18 | 171.82 | 73.07 | 0.74506 | 10 | Constant current mode
118 | 11.31 | 15.18 | 171.77 | 73.54 | 0.74506 | 11 | Constant current mode with 15 A reference
119 | 11.31 | 15.19 | 171.94 | 74.02 | 0.74457 | 12 | OVT fault
120 | 11.31 | 15.18 | 171.89 | 74.36 | 0.74506 | 13 | Fault cleared
121 | 11.31 | 15.18 | 171.89 | 75.08 | 0.74506
122 | 11.31 | 15.19 | 171.94 | 75.29 | 0.74457
123 | 11.31 | 15.18 | 171.87 | 75.73 | 0.74506
124 | 11.32 | 15.18 | 172.03 | 75.92 | 0.74572
125 | 11.32 | 15.18 | 172.01 | 76.37 | 0.74572
126 | 11.31 | 15.18 | 171.88 | 76.51 | 0.74506
127 | 11.32 | 15.19 | 172.05 | 76.9 | 0.74523
128 | 11.32 | 15.18 | 171.98 | 77 | 0.74572
129 | 11.32 | 15.18 | 172.03 | 77.22 | 0.74572
130 | 11.33 | 15.18 | 172.13 | 77.3 | 0.74638
131 | 11.31 | 15.18 | 171.88 | 77.4 | 0.74506
132 | 11.32 | 15.19 | 172.05 | 77.95 | 0.74523
133 | 11.31 | 15.18 | 171.89 | 78.11 | 0.74506
134 | 11.31 | 15.18 | 171.91 | 78.15 | 0.74506
135 | 11.32 | 15.19 | 172.04 | 78.23 | 0.74523
136 | 11.32 | 15.18 | 171.96 | 78.55 | 0.74572
137 | 11.31 | 15.19 | 171.94 | 78.54 | 0.74457
138 | 11.32 | 15.18 | 172 | 78.83 | 0.74572
139 | 11.32 | 15.18 | 171.99 | 79 | 0.74572
140 | 11.32 | 15.19 | 172.05 | 78.85 | 0.74523
141 | 11.32 | 15.18 | 171.96 | 78.93 | 0.74572
142 | 11.32 | 15.18 | 172.03 | 78.64 | 0.74572
143 | 11.32 | 15.18 | 172.01 | 78.63 | 0.74572
144 | 11.32 | 15.18 | 171.97 | 78.44 | 0.74572
145 | 11.32 | 15.19 | 172.04 | 78.35 | 0.74523
146 | 11.32 | 15.18 | 171.97 | 78.29 | 0.74572
147 | 11.32 | 15.18 | 172.01 | 78.13 | 0.74572
148 | 11.32 | 15.18 | 172.03 | 78.11 | 0.74572
149 | 11.32 | 15.18 | 171.96 | 77.79 | 0.74572
150 | 11.32 | 15.19 | 172.05 | 77.82 | 0.74523
151 | 11.31 | 15.18 | 171.89 | 77.59 | 0.74506
152 | 11.32 | 15.18 | 171.99 | 77.62 | 0.74572
153 | 11.32 | 15.19 | 172.05 | 77.34 | 0.74523
154 | 11.32 | 15.18 | 171.96 | 77.1 | 0.74572
155 | 11.32 | 15.19 | 172.03 | 77.26 | 0.74523
156 | 11.33 | 15.18 | 172.11 | 76.84 | 0.74638
157 | 11.33 | 15.18 | 172.08 | 76.93 | 0.74638
158 | 11.32 | 15.19 | 172.04 | 76.77 | 0.74523
159 | 11.73 | 5.93 | 69.68 | 77.53 | 1.9781
160 | 11.73 | 5.93 | 69.68 | 77.2 | 1.9781
161 | 11.73 | 5.93 | 69.68 | 77.11 | 1.9781
162 | 11.73 | 5.93 | 69.68 | 77.04 | 1.9781
163 | 11.73 | 5.93 | 69.67 | 76.94 | 1.9781
164 | 11.73 | 5.94 | 69.71 | 76.73 | 1.9747
165 | 11.73 | 5.94 | 69.69 | 76.5 | 1.9747
166 | 11.73 | 5.94 | 69.69 | 76.32 | 1.9747
167 | 11.73 | 5.94 | 69.69 | 75.87 | 1.9747
168 | 11.73 | 5.94 | 69.73 | 75.43 | 1.9747
169 | 11.73 | 5.94 | 69.69 | 75.22 | 1.9747
170 | 11.73 | 5.94 | 69.74 | 74.74 | 1.9747
171 | 11.73 | 5.94 | 69.74 | 74.36 | 1.9747
172 | 11.73 | 5.93 | 69.68 | 73.69 | 1.9781
173 | 11.73 | 5.94 | 69.72 | 73.13 | 1.9747
174 | 11.73 | 5.94 | 69.7 | 72.6 | 1.9747
175 | 11.73 | 5.94 | 69.69 | 72.04 | 1.9747
176 | 11.73 | 5.94 | 69.71 | 71.71 | 1.9747
177 | 11.73 | 5.93 | 69.68 | 71.36 | 1.9781
178 | 11.73 | 5.94 | 69.74 | 70.75 | 1.9747
179 | 11.73 | 5.94 | 69.74 | 70.25 | 1.9747
180 | 11.73 | 5.94 | 69.74 | 69.72 | 1.9747
181 | 11.73 | 5.93 | 69.68 | 69.31 | 1.9781
182 | 11.73 | 5.93 | 69.68 | 68.74 | 1.9781
183 | 11.73 | 5.93 | 69.7 | 68.34 | 1.9781
184 | 11.73 | 5.93 | 69.71 | 67.85 | 1.9781
185 | 11.73 | 5.93 | 69.72 | 67.34 | 1.9781
186 | 11.73 | 5.94 | 69.77 | 66.93 | 1.9747
187 | 11.73 | 5.94 | 69.77 | 66.54 | 1.9747
188 | 11.73 | 5.93 | 69.68 | 65.95 | 1.9781
189 | 11.73 | 5.93 | 69.69 | 65.65 | 1.9781
190 | 11.73 | 5.94 | 69.78 | 65.18 | 1.9747
191 | 11.73 | 5.94 | 69.77 | 64.85 | 1.9747
192 | 11.74 | 5.94 | 69.81 | 64.45 | 1.9764
193 | 11.73 | 5.94 | 69.79 | 63.9 | 1.9747
194 | 11.73 | 5.94 | 69.74 | 63.71 | 1.9747
195 | 11.73 | 5.94 | 69.72 | 63.32 | 1.9747
196 | 11.73 | 5.94 | 69.76 | 63.03 | 1.9747
197 | 11.73 | 5.94 | 69.73 | 62.68 | 1.9747
198 | 11.73 | 5.94 | 69.72 | 62.34 | 1.9747
199 | 11.73 | 5.94 | 69.76 | 62.05 | 1.9747
200 | 11.73 | 5.94 | 69.7 | 61.73 | 1.9747
201 | 11.73 | 5.94 | 69.72 | 61.36 | 1.9747
202 | 11.73 | 5.94 | 69.77 | 60.21 | 1.9747
203 | 11.73 | 5.94 | 69.76 | 56.07 | 1.9747
204 | 11.73 | 5.94 | 69.76 | 48.97 | 1.9747
205 | 11.73 | 5.94 | 69.76 | 48.03 | 1.9747
206 | 11.73 | 5.94 | 69.76 | 47.92 | 1.9747
207 | 11.73 | 5.94 | 69.75 | 47.83 | 1.9747
208 | 11.73 | 5.94 | 69.77 | 47.72 | 1.9747
209 | 9.85 | 5.22 | 51.42 | 47.63 | 1.887
210 | 7.82 | 3.96 | 30.98 | 47.55 | 1.9747
211 | 5.92 | 3.19 | 18.96 | 47.53 | 1.8558
212 | 5.86 | 2.96 | 17.41 | 47.46 | 1.9797
213 | 5.86 | 2.96 | 17.42 | 47.39 | 1.9797
214 | 5.86 | 2.97 | 17.41 | 47.32 | 1.9731
215 | 5.86 | 2.96 | 17.41 | 47.25 | 1.9797
216 | 5.86 | 2.97 | 17.43 | 47.17 | 1.9731
217 | 5.86 | 2.97 | 17.41 | 47.07 | 1.9731
218 | 5.86 | 2.96 | 17.41 | 46.96 | 1.9797
219 | 5.86 | 2.96 | 17.4 | 46.84 | 1.9797
220 | 5.86 | 2.97 | 17.43 | 46.72 | 1.9731
221 | 5.86 | 2.96 | 17.42 | 46.63 | 1.9797
222 | 6.84 | 3.46 | 23.71 | 46.5 | 1.9769
223 | 7.82 | 3.95 | 30.96 | 46.35 | 1.9797
224 | 8.79 | 4.45 | 39.18 | 46.24 | 1.9753
225 | 10.73 | 5.37 | 57.73 | 46.1 | 1.9981
226 | 12.71 | 6.43 | 81.75 | 45.95 | 1.9767
227 | 14.66 | 7.42 | 108.85 | 45.83 | 1.9757
228 | 18.33 | 8.94 | 164 | 45.7 | 2.0503
229 | 18.58 | 9.4 | 174.85 | 45.59 | 1.9766
230 | 19.46 | 9.59 | 186.76 | 45.48 | 2.0292
231 | 19.56 | 9.9 | 193.68 | 45.39 | 1.9758
232 | 19.55 | 9.9 | 193.66 | 45.29 | 1.9747
233 | 19.55 | 9.9 | 193.62 | 45.22 | 1.9747
234 | 19.56 | 9.9 | 193.79 | 45.2 | 1.9758
235 | 19.55 | 9.89 | 193.6 | 45.14 | 1.9767
236 | 19.56 | 9.9 | 193.68 | 45.14 | 1.9758
237 | 15.69 | 8.06 | 126.63 | 45.15 | 1.9467
238 | 12.26 | 6.43 | 78.94 | 45.19 | 1.9067
239 | 11.73 | 5.94 | 69.76 | 45.25 | 1.9747
240 | 11.73 | 5.94 | 69.75 | 45.26 | 1.9747
241 | 11.73 | 5.94 | 69.71 | 45.34 | 1.9747
242 | 11.73 | 5.94 | 69.75 | 45.37 | 1.9747
243 | 11.73 | 5.94 | 69.72 | 45.39 | 1.9747
244 | 11.73 | 5.94 | 69.74 | 45.45 | 1.9747
245 | 11.73 | 5.94 | 69.74 | 45.5 | 1.9747
246 | 11.73 | 5.94 | 69.74 | 45.5 | 1.9747
247 | 11.65 | 7.86 | 91.68 | 45.47 | 1.4822
248 | 11.65 | 7.86 | 91.74 | 45.47 | 1.4822
249 | 11.65 | 7.86 | 91.67 | 45.48 | 1.4822
250 | 11.65 | 7.86 | 91.72 | 45.44 | 1.4822
251 | 11.65 | 7.86 | 91.75 | 45.42 | 1.4822
252 | 11.65 | 7.86 | 91.67 | 45.4 | 1.4822
253 | 11.65 | 7.86 | 91.72 | 45.38 | 1.4822
254 | 11.64 | 7.86 | 91.59 | 45.36 | 1.4809
255 | 12.61 | 8.51 | 107.51 | 45.32 | 1.4818
256 | 13.6 | 9.17 | 124.78 | 45.3 | 1.4831
257 | 13.59 | 9.17 | 124.74 | 45.31 | 1.482
258 | 14.56 | 9.82 | 143.15 | 45.29 | 1.4827
259 | 15.54 | 10.48 | 163 | 45.31 | 1.4828
260 | 17.46 | 11.78 | 205.92 | 45.25 | 1.4822
261 | 17.48 | 11.79 | 206.21 | 45.24 | 1.4826
262 | 17.47 | 11.8 | 206.21 | 45.25 | 1.4805
263 | 17.47 | 11.79 | 206.14 | 45.25 | 1.4818
264 | 17.48 | 11.8 | 206.3 | 45.31 | 1.4814
265 | 17.47 | 11.79 | 206.18 | 45.33 | 1.4818
266 | 17.48 | 11.79 | 206.21 | 45.39 | 1.4826
267 | 17.47 | 11.8 | 206.22 | 45.43 | 1.4805
268 | 12.63 | 8.53 | 107.87 | 45.47 | 1.4807
269 | 11.65 | 7.86 | 91.73 | 45.6 | 1.4822
270 | 11.65 | 7.86 | 91.72 | 45.68 | 1.4822
271 | 11.65 | 7.86 | 91.65 | 45.77 | 1.4822
272 | 11.65 | 7.86 | 91.66 | 45.87 | 1.4822
273 | 11.65 | 7.86 | 91.67 | 45.91 | 1.4822
274 | 11.65 | 7.86 | 91.66 | 45.97 | 1.4822
275 | 11.65 | 7.86 | 91.72 | 46.05 | 1.4822
276 | 11.65 | 7.86 | 91.68 | 46.12 | 1.4822
277 | 11.65 | 7.86 | 91.64 | 46.15 | 1.4822
278 | 11.65 | 7.86 | 91.68 | 46.15 | 1.4822
279 | 11.65 | 7.86 | 91.74 | 46.24 | 1.4822
280 | 11.65 | 7.86 | 91.72 | 46.27 | 1.4822
281 | 11.65 | 7.86 | 91.74 | 46.31 | 1.4822
282 | 11.65 | 7.86 | 91.74 | 46.32 | 1.4822
283 | 11.65 | 7.86 | 91.66 | 46.32 | 1.4822
284 | 11.65 | 7.86 | 91.74 | 46.31 | 1.4822
285 | 11.65 | 7.86 | 91.74 | 46.3 | 1.4822
286 | 11.65 | 7.86 | 91.67 | 46.29 | 1.4822
287 | 11.65 | 7.86 | 91.75 | 46.27 | 1.4822
288 | 11.65 | 7.86 | 91.73 | 46.27 | 1.4822
289 | 11.65 | 7.86 | 91.72 | 46.26 | 1.4822
290 | 11.65 | 7.87 | 91.7 | 46.22 | 1.4803
291 | 11.65 | 7.86 | 91.69 | 46.2 | 1.4822
292 | 11.65 | 7.86 | 91.72 | 46.19 | 1.4822
293 | 11.65 | 7.86 | 91.7 | 46.16 | 1.4822
294 | 11.65 | 7.86 | 91.67 | 46.16 | 1.4822
295 | 11.65 | 7.86 | 91.74 | 46.13 | 1.4822
296 | 11.65 | 7.86 | 91.7 | 46.09 | 1.4822
297 | 11.66 | 7.86 | 91.77 | 46.08 | 1.4835
298 | 11.65 | 7.86 | 91.74 | 46.04 | 1.4822
299 | 11.65 | 7.86 | 91.68 | 46.04 | 1.4822
300 | 11.65 | 7.86 | 91.67 | 45.99 | 1.4822
301 | 11.65 | 7.87 | 91.75 | 45.99 | 1.4803
302 | 11.65 | 7.86 | 91.69 | 45.94 | 1.4822
303 | 11.65 | 7.86 | 91.72 | 45.95 | 1.4822
304 | 11.65 | 7.86 | 91.75 | 45.9 | 1.4822
305 | 11.66 | 7.87 | 91.81 | 45.88 | 1.4816
306 | 11.65 | 7.86 | 91.68 | 45.85 | 1.4822
307 | 11.65 | 7.86 | 91.75 | 45.83 | 1.4822
308 | 11.65 | 7.87 | 91.76 | 45.79 | 1.4803
309 | 11.65 | 7.86 | 91.74 | 45.78 | 1.4822
310 | 11.65 | 7.87 | 91.76 | 45.77 | 1.4803
311 | 11.65 | 7.87 | 91.76 | 45.73 | 1.4803
312 | 11.66 | 7.86 | 91.78 | 45.72 | 1.4835
313 | 11.65 | 7.87 | 91.77 | 45.66 | 1.4803
314 | 11.65 | 7.86 | 91.75 | 45.67 | 1.4822
315 | 11.66 | 7.86 | 91.79 | 45.62 | 1.4835
316 | 11.65 | 7.87 | 91.78 | 45.61 | 1.4803
317 | 11.65 | 7.86 | 91.7 | 45.57 | 1.4822
318 | 11.65 | 7.86 | 91.73 | 45.54 | 1.4822
319 | 11.65 | 7.87 | 91.76 | 45.52 | 1.4803
320 | 11.81 | 4.29 | 50.7 | 45.57 | 2.7529
321 | 11.81 | 4.28 | 50.63 | 45.54 | 2.7593
322 | 11.81 | 4.27 | 50.55 | 45.5 | 2.7658
323 | 11.8 | 4.28 | 50.6 | 45.47 | 2.757
324 | 11.81 | 4.27 | 50.51 | 45.45 | 2.7658
325 | 11.81 | 4.28 | 50.61 | 45.42 | 2.7593
326 | 11.81 | 4.29 | 50.69 | 45.36 | 2.7529
327 | 11.81 | 4.28 | 50.6 | 45.34 | 2.7593
328 | 11.81 | 4.28 | 50.56 | 45.3 | 2.7593
329 | 11.8 | 4.28 | 50.54 | 45.23 | 2.757
330 | 9.76 | 5.18 | 50.59 | 45.16 | 1.8842
331 | 6.67 | 7.59 | 50.7 | 45.06 | 0.87879
332 | 4.51 | 11.2 | 50.6 | 44.99 | 0.40268
333 | 4.51 | 11.21 | 50.65 | 44.96 | 0.40232
334 | 4.5 | 11.22 | 50.61 | 44.91 | 0.40107
335 | 4.5 | 11.22 | 50.62 | 44.86 | 0.40107
336 | 3.34 | 15.1 | 50.58 | 44.77 | 0.22119
337 | 3.34 | 15.11 | 50.62 | 44.72 | 0.22105
338 | 3.34 | 15.13 | 50.56 | 44.65 | 0.22075
339 | 3.34 | 15.12 | 50.65 | 44.58 | 0.2209
340 | 3.35 | 15.11 | 50.72 | 44.57 | 0.22171
341 | 3.34 | 15.1 | 50.59 | 44.5 | 0.22119
342 | 3.35 | 15.09 | 50.63 | 44.47 | 0.222
343 | 3.34 | 15.1 | 50.57 | 44.43 | 0.22119
344 | 3.33 | 15.16 | 50.57 | 44.39 | 0.21966
345 | 3.34 | 15.15 | 50.75 | 44.37 | 0.22046
346 | 3.34 | 15.18 | 50.73 | 44.35 | 0.22003
347 | 3.32 | 15.19 | 50.57 | 44.32 | 0.21856
348 | 4.5 | 11.26 | 50.71 | 44.32 | 0.39964
349 | 8.71 | 5.84 | 50.92 | 44.35 | 1.4914
350 | 11.8 | 4.29 | 50.71 | 44.36 | 2.7506
351 | 12.81 | 3.94 | 50.59 | 44.37 | 3.2513
352 | 13.83 | 3.66 | 50.65 | 44.33 | 3.7787
353 | 13.82 | 3.66 | 50.63 | 44.32 | 3.776
354 | 13.82 | 3.66 | 50.61 | 44.31 | 3.776
355 | 13.83 | 3.66 | 50.65 | 44.31 | 3.7787
356 | 13.82 | 3.66 | 50.64 | 44.31 | 3.776
357 | 13.83 | 3.65 | 50.57 | 44.34 | 3.789
358 | 13.83 | 3.66 | 50.63 | 44.36 | 3.7787
359 | 13.82 | 3.65 | 50.58 | 44.39 | 3.7863
360 | 13.82 | 3.66 | 50.61 | 44.42 | 3.776
361 | 13.83 | 3.66 | 50.68 | 44.44 | 3.7787
362 | 13.82 | 3.65 | 50.59 | 44.5 | 3.7863
363 | 13.83 | 3.66 | 50.66 | 44.51 | 3.7787
364 | 13.83 | 3.65 | 50.57 | 44.56 | 3.789
365 | 13.83 | 3.65 | 50.61 | 44.57 | 3.789
366 | 13.8 | 3.66 | 50.56 | 44.63 | 3.7705
367 | 13.83 | 3.66 | 50.66 | 44.67 | 3.7787
368 | 13.83 | 3.65 | 50.49 | 44.7 | 3.789
369 | 13.83 | 3.66 | 50.74 | 44.78 | 3.7787
370 | 13.83 | 3.65 | 50.53 | 44.85 | 3.789
371 | 13.83 | 3.65 | 50.59 | 44.94 | 3.789
372 | 13.83 | 3.66 | 50.68 | 44.97 | 3.7787
373 | 13.83 | 3.65 | 50.58 | 45 | 3.789
374 | 13.83 | 3.65 | 50.56 | 45.05 | 3.789
375 | 13.82 | 3.66 | 50.73 | 45.13 | 3.776
376 | 13.83 | 3.66 | 50.65 | 45.19 | 3.7787
377 | 13.66 | 7.36 | 100.59 | 45.2 | 1.856
378 | 13.67 | 7.43 | 101.6 | 45.22 | 1.8398
379 | 13.67 | 7.32 | 100.22 | 45.23 | 1.8675
380 | 13.67 | 7.45 | 101.89 | 45.27 | 1.8349
381 | 13.67 | 7.4 | 101.28 | 45.3 | 1.8473
382 | 13.66 | 7.36 | 100.63 | 45.4 | 1.856
383 | 13.68 | 7.46 | 102.13 | 45.44 | 1.8338
384 | 13.67 | 7.47 | 102.27 | 45.51 | 1.83
385 | 13.67 | 7.33 | 100.29 | 45.58 | 1.8649
386 | 13.66 | 7.48 | 102.29 | 45.68 | 1.8262
387 | 13.66 | 7.33 | 100.19 | 45.75 | 1.8636
388 | 13.67 | 7.34 | 100.44 | 45.83 | 1.8624
389 | 13.66 | 7.38 | 100.91 | 45.92 | 1.8509
390 | 13.66 | 7.48 | 102.26 | 45.96 | 1.8262
391 | 13.67 | 7.34 | 100.41 | 46.03 | 1.8624
392 | 13.66 | 7.43 | 101.64 | 46.16 | 1.8385
393 | 13.66 | 7.41 | 101.33 | 46.25 | 1.8435
394 | 13.67 | 7.48 | 102.32 | 46.38 | 1.8275
395 | 13.67 | 7.46 | 102.06 | 46.5 | 1.8324
396 | 13.67 | 7.46 | 102.05 | 46.6 | 1.8324
397 | 13.66 | 7.32 | 100.15 | 46.69 | 1.8661
398 | 13.66 | 7.44 | 101.68 | 46.8 | 1.836
399 | 13.66 | 7.47 | 102.11 | 46.89 | 1.8286
400 | 13.66 | 7.33 | 100.19 | 46.96 | 1.8636
401 | 13.67 | 7.39 | 101.06 | 47.07 | 1.8498
402 | 13.67 | 7.48 | 102.32 | 47.18 | 1.8275
403 | 13.66 | 7.32 | 100.12 | 47.28 | 1.8661
404 | 13.66 | 7.37 | 100.73 | 47.36 | 1.8535
405 | 13.66 | 7.45 | 101.82 | 47.46 | 1.8336
406 | 13.66 | 7.41 | 101.34 | 47.53 | 1.8435
407 | 13.66 | 7.45 | 101.84 | 47.66 | 1.8336
408 | 13.66 | 7.41 | 101.37 | 47.73 | 1.8435
409 | 13.67 | 7.33 | 100.27 | 47.8 | 1.8649
410 | 13.66 | 7.46 | 102.04 | 47.91 | 1.8311
411 | 13.67 | 7.38 | 100.99 | 48.02 | 1.8523
412 | 13.67 | 7.46 | 102.08 | 48.75 | 1.8324
413 | 13.67 | 7.35 | 100.6 | 56.36 | 1.8599
414 | 13.66 | 7.34 | 100.35 | 57.59 | 1.861
415 | 13.67 | 7.47 | 102.14 | 61.25 | 1.83
416 | 13.67 | 7.38 | 100.97 | 61.5 | 1.8523
417 | 13.66 | 7.44 | 101.75 | 61.74 | 1.836
418 | 13.66 | 7.36 | 100.68 | 62.02 | 1.856
419 | 13.67 | 7.4 | 101.18 | 62.26 | 1.8473
420 | 13.67 | 7.44 | 101.83 | 62.58 | 1.8374
421 | 13.67 | 7.43 | 101.68 | 62.86 | 1.8398
422 | 13.66 | 7.37 | 100.83 | 63.07 | 1.8535
423 | 13.66 | 7.39 | 101.09 | 63.33 | 1.8484
424 | 13.66 | 7.44 | 101.74 | 63.76 | 1.836
425 | 13.66 | 7.42 | 101.45 | 63.99 | 1.841
426 | 13.66 | 7.44 | 101.69 | 64.19 | 1.836
427 | 13.67 | 7.39 | 101.07 | 64.54 | 1.8498
428 | 13.67 | 7.41 | 101.36 | 64.89 | 1.8448
429 | 13.67 | 7.42 | 101.49 | 65.22 | 1.8423
430 | 13.66 | 7.39 | 101.02 | 65.68 | 1.8484
431 | 13.66 | 7.4 | 101.25 | 65.82 | 1.8459
432 | 13.66 | 7.42 | 101.47 | 66.12 | 1.841
433 | 13.66 | 7.41 | 101.3 | 66.64 | 1.8435
434 | 13.66 | 7.42 | 101.41 | 67.04 | 1.841
435 | 15.71 | 6.46 | 101.64 | 67.18 | 2.4319
436 | 16.72 | 6.06 | 101.43 | 67.54 | 2.7591
437 | 17.74 | 5.7 | 101.14 | 68.04 | 3.1123
438 | 18.74 | 5.39 | 101.2 | 68.47 | 3.4768
439 | 19.68 | 5.25 | 103.38 | 68.8 | 3.7486
440 | 21.78 | 4.65 | 101.31 | 69.09 | 4.6839
441 | 21.78 | 4.64 | 101.26 | 69.38 | 4.694
442 | 21.78 | 4.64 | 101.21 | 69.72 | 4.694
443 | 21.78 | 4.64 | 101.15 | 69.83 | 4.694
444 | 21.78 | 4.64 | 101.18 | 69.97 | 4.694
445 | 21.78 | 4.64 | 101.27 | 69.47 | 4.694
446 | 21.78 | 4.64 | 101.22 | 69.7 | 4.694
447 | 21.78 | 4.64 | 101.19 | 69.78 | 4.694
448 | 21.78 | 4.65 | 101.31 | 70.03 | 4.6839
449 | 21.77 | 4.64 | 101.19 | 70.31 | 4.6918
450 | 18.76 | 5.4 | 101.42 | 70.61 | 3.4741
451 | 14.69 | 6.88 | 101.22 | 70.69 | 2.1352
452 | 12.64 | 8.02 | 101.43 | 70.75 | 1.5761
453 | 12.64 | 8.02 | 101.5 | 70.82 | 1.5761
454 | 11.61 | 8.72 | 101.26 | 71.12 | 1.3314
455 | 11.6 | 8.71 | 101.14 | 71.21 | 1.3318
456 | 11.61 | 8.71 | 101.23 | 71.15 | 1.333
457 | 11.61 | 8.71 | 101.24 | 71.42 | 1.333
458 | 11.61 | 8.71 | 101.21 | 71.56 | 1.333
459 | 11.61 | 8.72 | 101.33 | 71.66 | 1.3314
460 | 11.61 | 8.71 | 101.23 | 71.69 | 1.333
461 | 11.61 | 8.71 | 101.3 | 71.84 | 1.333
462 | 11.61 | 8.72 | 101.34 | 71.71 | 1.3314
463 | 11.61 | 8.71 | 101.2 | 71.63 | 1.333
464 | 11.61 | 8.71 | 101.23 | 71.62 | 1.333
465 | 11.61 | 8.71 | 101.23 | 71.38 | 1.333
466 | 11.61 | 8.71 | 101.21 | 71.49 | 1.333
467 | 11.61 | 8.71 | 101.24 | 71.46 | 1.333
468 | 11.33 | 15.19 | 172.15 | 70.67 | 0.74589
469 | 11.33 | 15.18 | 172.07 | 70.72 | 0.74638
470 | 11.33 | 15.18 | 172.13 | 70.47 | 0.74638
471 | 11.33 | 15.18 | 172.12 | 70.48 | 0.74638
472 | 11.33 | 15.18 | 172.08 | 70.5 | 0.74638
473 | 11.33 | 15.19 | 172.14 | 70.16 | 0.74589
474 | 11.31 | 15.18 | 171.78 | 69.95 | 0.74506
475 | 14.26 | 15.18 | 216.68 | 69.84 | 0.93939
476 | 16.31 | 15.18 | 247.82 | 69.72 | 1.0744
477 | 16.32 | 15.18 | 247.82 | 69.67 | 1.0751
478 | 18.31 | 15.19 | 278.22 | 69.72 | 1.2054
479 | 19.31 | 15.18 | 293.34 | 69.63 | 1.2721
480 | 19.31 | 15.18 | 293.34 | 69.75 | 1.2721
481 | 19.31 | 15.19 | 293.42 | 69.77 | 1.2712
482 | 19.31 | 15.18 | 293.3 | 69.82 | 1.2721
483 | 19.3 | 15.19 | 293.31 | 69.88 | 1.2706
484 | 19.3 | 15.18 | 293.27 | 70.11 | 1.2714
485 | 20.31 | 15.18 | 308.51 | 70.18 | 1.3379
486 | 20.31 | 15.19 | 308.64 | 70.39 | 1.3371
487 | 21.31 | 15.19 | 323.97 | 70.75 | 1.4029
488 | 21.31 | 15.19 | 324.02 | 71.16 | 1.4029
489 | 21.32 | 15.2 | 324.12 | 71.63 | 1.4026
490 | 21.31 | 15.19 | 323.91 | 72.08 | 1.4029
491 | 21.31 | 15.2 | 324.05 | 72.58 | 1.402
492 | 21.32 | 15.19 | 324.02 | 73.17 | 1.4036
493 | 21.31 | 15.19 | 323.94 | 73.61 | 1.4029
494 | 21.31 | 15.2 | 324.07 | 73.96 | 1.402
495 | 21.31 | 15.19 | 323.93 | 74.46 | 1.4029
496 | 21.32 | 15.19 | 324.12 | 75.04 | 1.4036
497 | 21.31 | 15.19 | 323.92 | 75.64 | 1.4029
498 | 21.31 | 15.19 | 323.88 | 76.25 | 1.4029
499 | 21.32 | 15.2 | 324.16 | 76.75 | 1.4026
500 | 21.31 | 15.19 | 323.87 | 77.36 | 1.4029
501 | 21.31 | 15.19 | 324 | 77.91 | 1.4029
502 | 21.32 | 15.19 | 324.09 | 78.61 | 1.4036
503 | 21.31 | 15.19 | 323.89 | 78.99 | 1.4029
504 | 21.32 | 15.2 | 324.14 | 79.67 | 1.4026
505 | 21.31 | 15.19 | 323.8 | 80.51 | 1.4029
506 | 21.32 | 15.19 | 324.08 | 80.89 | 1.4036
507 | 21.31 | 15.19 | 324.01 | 81.58 | 1.4029
508 | 21.31 | 15.19 | 323.89 | 82.03 | 1.4029
509 | 21.31 | 15.2 | 324.03 | 82.73 | 1.402
510 | 21.31 | 15.19 | 323.96 | 83.56 | 1.4029
511 | 21.32 | 15.19 | 324.05 | 83.88 | 1.4036
512 | 21.31 | 15.2 | 324.04 | 84.23 | 1.402
513 | 21.32 | 15.19 | 323.99 | 84.73 | 1.4036
514 | 21.32 | 15.19 | 324.11 | 85.34 | 1.4036
515 | 21.31 | 15.19 | 323.96 | 85.64 | 1.4029
516 | 21.32 | 15.19 | 324.02 | 86.34 | 1.4036
517 | 21.31 | 15.2 | 324.04 | 86.62 | 1.402
518 | 21.32 | 15.19 | 323.99 | 87.34 | 1.4036
519 | 21.31 | 15.19 | 323.99 | 87.64 | 1.4029
520 | 21.31 | 15.19 | 323.98 | 88.18 | 1.4029
521 | 21.31 | 15.19 | 323.91 | 88.88 | 1.4029
522 | 21.31 | 15.2 | 324.04 | 89.28 | 1.402
523 | 21.32 | 15.19 | 324.03 | 89.79 | 1.4036
524 | 21.31 | 15.19 | 323.95 | 90.27 | 1.4029
525 | 21.31 | 15.21 | 324.15 | 90.89 | 1.4011
526 | 21.31 | 15.2 | 323.93 | 91.4 | 1.402
527 | 21.31 | 15.2 | 324.05 | 91.53 | 1.402
528 | 21.31 | 15.2 | 323.99 | 91.85 | 1.402
529 | 21.31 | 15.2 | 323.93 | 92.6 | 1.402
530 | 21.31 | 15.2 | 324.04 | 92.86 | 1.402
531 | 21.31 | 15.19 | 324.01 | 93.26 | 1.4029
532 | 21.31 | 15.2 | 324.03 | 93.47 | 1.402
533 | 21.31 | 15.2 | 323.99 | 93.74 | 1.402
534 | 21.31 | 15.2 | 323.92 | 94.22 | 1.402
535 | 21.31 | 15.2 | 324.02 | 94.7 | 1.402
536 | 21.31 | 15.2 | 324.02 | 94.68 | 1.402
537 | 21.31 | 15.2 | 323.98 | 95.21 | 1.402
538 | 21.31 | 15.2 | 324.12 | 95.65 | 1.402
539 | 21.31 | 15.19 | 323.87 | 95.98 | 1.4029
540 | 21.31 | 15.2 | 324.12 | 96.23 | 1.402
541 | 21.31 | 15.19 | 324.02 | 96.58 | 1.4029
542 | 21.31 | 15.2 | 324.04 | 96.85 | 1.402
543 | 21.31 | 15.2 | 324.01 | 96.86 | 1.402
544 | 21.31 | 15.19 | 323.97 | 96.94 | 1.4029
545 | 21.31 | 15.2 | 324.11 | 97.31 | 1.402
546 | 21.31 | 15.2 | 323.96 | 97.67 | 1.402
547 | 21.31 | 15.19 | 324.01 | 97.6 | 1.4029
548 | 21.31 | 15.2 | 324.04 | 97.85 | 1.402
549 | 21.31 | 15.19 | 323.87 | 98.05 | 1.4029
550 | 21.31 | 15.2 | 323.98 | 98.24 | 1.402
551 | 21.31 | 15.2 | 323.98 | 98.23 | 1.402
552 | 21.31 | 15.19 | 323.88 | 98.87 | 1.4029
553 | 21.31 | 15.2 | 324.05 | 99.01 | 1.402
554 | 21.31 | 15.19 | 323.9 | 98.85 | 1.4029
555 | 21.31 | 15.2 | 324.06 | 99.27 | 1.402
556 | 21.31 | 15.2 | 324.09 | 99.39 | 1.402
557 | 21.31 | 15.19 | 323.9 | 99.57 | 1.4029
558 | 21.31 | 15.2 | 324.02 | 100 | 1.402
559 | 21.32 | 15.2 | 324.13 | 100.23 | 1.4026
560 | 21.31 | 15.2 | 323.92 | 100.41 | 1.402
561 | 21.31 | 15.2 | 323.96 | 100.53 | 1.402
562 | 21.31 | 15.19 | 323.93 | 100.52 | 1.4029
563 | 21.31 | 15.2 | 324 | 100.73 | 1.402
564 | 21.31 | 15.19 | 323.98 | 100.98 | 1.4029
565 | 21.31 | 15.19 | 324.01 | 101.07 | 1.4029
566 | 21.31 | 15.2 | 323.96 | 101.09 | 1.402
567 | 21.31 | 15.19 | 323.96 | 101.13 | 1.4029
568 | 21.31 | 15.2 | 324.02 | 101.37 | 1.402
569 | 21.31 | 15.19 | 323.98 | 101.45 | 1.4029
570 | 21.31 | 15.2 | 323.93 | 102.21 | 1.402
571 | 21.31 | 15.2 | 324.04 | 102.18 | 1.402
572 | 21.31 | 15.19 | 323.88 | 102.45 | 1.4029
573 | 21.31 | 15.2 | 323.99 | 102.27 | 1.402
574 | 21.31 | 15.19 | 323.86 | 102.41 | 1.4029
575 | 21.31 | 15.2 | 323.96 | 102.49 | 1.402
576 | 21.31 | 15.2 | 324.05 | 102.4 | 1.402
577 | 21.31 | 15.2 | 324.03 | 102.91 | 1.402
578 | 21.31 | 15.2 | 324.04 | 102.77 | 1.402
579 | 21.31 | 15.19 | 323.87 | 103.15 | 1.4029
580 | 21.31 | 15.2 | 324.03 | 103.02 | 1.402
581 | 21.31 | 15.2 | 324.06 | 103.46 | 1.402
582 | 21.31 | 15.19 | 324.01 | 103.3 | 1.4029
583 | 21.3 | 15.2 | 323.94 | 103.56 | 1.4013
584 | 21.31 | 15.19 | 323.88 | 104.09 | 1.4029
585 | 21.31 | 15.2 | 324.03 | 104.32 | 1.402
586 | 21.31 | 15.2 | 323.93 | 104.53 | 1.402
587 | 21.31 | 15.2 | 323.94 | 104.76 | 1.402
588 | 21.31 | 15.2 | 324 | 104.58 | 1.402
589 | 21.31 | 15.19 | 323.87 | 105.11 | 1.4029
590 | 21.31 | 15.2 | 324.04 | 105.1 | 1.402
591 | 21.31 | 15.2 | 324.03 | 105.01 | 1.402
592 | 21.31 | 15.2 | 324.04 | 105.11 | 1.402
593 | 21.31 | 15.2 | 324.07 | 105.42 | 1.402
594 | 21.3 | 15.2 | 323.84 | 105.67 | 1.4013
595 | 21.31 | 15.2 | 324.08 | 105.65 | 1.402
596 | 21.31 | 15.2 | 323.95 | 105.61 | 1.402
597 | 21.31 | 15.2 | 324.12 | 105.93 | 1.402
598 | 21.31 | 15.2 | 324.01 | 105.88 | 1.402
599 | 21.31 | 15.2 | 324.05 | 105.91 | 1.402
600 | 21.31 | 15.2 | 324.07 | 105.82 | 1.402
601 | 21.31 | 15.2 | 323.92 | 105.81 | 1.402
602 | 21.31 | 15.2 | 324.02 | 105.89 | 1.402
603 | 21.31 | 15.2 | 323.99 | 105.92 | 1.402
604 | 21.31 | 15.2 | 323.94 | 106.27 | 1.402
605 | 21.31 | 15.2 | 324.04 | 106.04 | 1.402
606 | 21.31 | 15.19 | 323.91 | 106.11 | 1.4029
607 | 21.31 | 15.2 | 324.01 | 106.05 | 1.402
608 | 21.31 | 15.2 | 324.05 | 106.31 | 1.402
609 | 21.31 | 15.2 | 323.93 | 105.95 | 1.402
610 | 21.31 | 15.2 | 324.11 | 105.75 | 1.402
611 | 21.31 | 15.19 | 323.9 | 105.91 | 1.4029
612 | 21.31 | 15.2 | 324.1 | 106.09 | 1.402
613 | 21.31 | 15.19 | 324.01 | 105.99 | 1.4029
614 | 21.31 | 15.19 | 324.01 | 106.24 | 1.4029
615 | 21.32 | 15.2 | 324.18 | 106.15 | 1.4026
616 | 21.31 | 15.19 | 323.89 | 106.4 | 1.4029
617 | 21.31 | 15.2 | 324.01 | 106.28 | 1.402
618 | 21.31 | 15.19 | 323.89 | 106.23 | 1.4029
619 | 21.31 | 15.2 | 323.96 | 106.47 | 1.402
620 | 21.31 | 15.2 | 324.07 | 106.31 | 1.402
621 | 21.31 | 15.19 | 324.01 | 106.44 | 1.4029
622 | 21.31 | 15.2 | 323.98 | 106.79 | 1.402
623 | 21.31 | 15.19 | 323.95 | 106.62 | 1.4029
624 | 21.31 | 15.2 | 324.06 | 106.78 | 1.402
625 | 21.31 | 15.2 | 324.05 | 106.47 | 1.402
626 | 21.31 | 15.19 | 323.98 | 106.57 | 1.4029
627 | 21.31 | 15.2 | 323.98 | 106.65 | 1.402
628 | 21.31 | 15.19 | 323.94 | 106.71 | 1.4029
629 | 21.31 | 15.2 | 324.06 | 106.58 | 1.402
630 | 21.31 | 15.19 | 324.01 | 106.88 | 1.4029
631 | 21.31 | 15.19 | 323.96 | 106.89 | 1.4029
632 | 21.32 | 15.2 | 324.16 | 107.14 | 1.4026
633 | 21.32 | 15.19 | 324.02 | 107.11 | 1.4036
634 | 21.31 | 15.2 | 324.06 | 107.29 | 1.402
635 | 21.31 | 15.19 | 323.99 | 107.11 | 1.4029
636 | 21.31 | 15.19 | 323.96 | 107.16 | 1.4029
637 | 21.31 | 15.2 | 323.94 | 106.98 | 1.402
638 | 21.31 | 15.19 | 323.78 | 106.94 | 1.4029
639 | 21.31 | 15.2 | 324.04 | 107 | 1.402
640 | 21.31 | 15.19 | 323.96 | 107.12 | 1.4029
641 | 21.31 | 15.19 | 323.96 | 107.18 | 1.4029
642 | 21.31 | 15.2 | 324.04 | 107.14 | 1.402
643 | 21.31 | 15.19 | 323.9 | 107.14 | 1.4029
644 | 21.31 | 15.2 | 324.04 | 106.99 | 1.402
645 | 21.31 | 15.19 | 323.92 | 107.21 | 1.4029
646 | 21.31 | 15.19 | 323.98 | 107.13 | 1.4029
647 | 21.32 | 15.19 | 324.1 | 107.17 | 1.4036
648 | 21.32 | 15.19 | 324 | 106.72 | 1.4036
649 | 21.32 | 15.2 | 324.15 | 106.74 | 1.4026
650 | 21.32 | 15.19 | 324 | 107.25 | 1.4036
651 | 21.31 | 15.19 | 323.99 | 106.83 | 1.4029
652 | 21.32 | 15.19 | 324.08 | 106.87 | 1.4036
653 | 21.32 | 15.19 | 323.98 | 106.66 | 1.4036
654 | 21.32 | 15.2 | 324.14 | 106.63 | 1.4026
655 | 21.31 | 15.19 | 323.89 | 106.83 | 1.4029
656 | 21.31 | 15.19 | 323.99 | 106.54 | 1.4029
657 | 21.32 | 15.19 | 324.06 | 106.56 | 1.4036
658 | 21.32 | 15.19 | 324.01 | 106.62 | 1.4036
659 | 21.31 | 15.19 | 324.01 | 106.29 | 1.4029
660 | 21.32 | 15.19 | 323.97 | 106.03 | 1.4036
661 | 21.32 | 15.19 | 324.1 | 106.47 | 1.4036
662 | 21.32 | 15.19 | 324.04 | 106.48 | 1.4036
663 | 21.32 | 15.19 | 324.01 | 106.58 | 1.4036
664 | 21.32 | 15.19 | 324.06 | 106.64 | 1.4036
665 | 21.32 | 15.19 | 323.96 | 106.56 | 1.4036
666 | 22.35 | 14.5 | 324.09 | 106.44 | 1.5414
667 | 24.39 | 13.27 | 323.97 | 106.74 | 1.838
668 | 25.41 | 12.74 | 323.93 | 106.87 | 1.9945
669 | 25.42 | 12.74 | 324.15 | 106.9 | 1.9953
670 | 25.41 | 12.74 | 323.92 | 106.46 | 1.9945
671 | 25.41 | 12.74 | 324.05 | 106.58 | 1.9945
672 | 25.41 | 12.74 | 323.95 | 106.6 | 1.9945
673 | 25.41 | 12.74 | 323.98 | 106.61 | 1.9945
674 | 25.41 | 12.74 | 324.01 | 106.65 | 1.9945
675 | 25.41 | 12.74 | 323.88 | 106.56 | 1.9945
676 | 25.41 | 12.74 | 324.03 | 106.51 | 1.9945
677 | 25.42 | 12.74 | 324 | 106.75 | 1.9953
678 | 25.41 | 12.74 | 323.97 | 106.66 | 1.9945
679 | 25.41 | 12.74 | 324.04 | 106.7 | 1.9945
680 | 26.43 | 12.25 | 323.9 | 106.65 | 2.1576
681 | 26.43 | 12.25 | 324.03 | 106.62 | 2.1576
682 | 27.31 | 11.86 | 323.95 | 106.77 | 2.3027
683 | 27.31 | 11.86 | 324.01 | 106.66 | 2.3027
684 | 27.31 | 11.86 | 324.02 | 106.83 | 2.3027
685 | 27.31 | 11.86 | 323.95 | 106.65 | 2.3027
686 | 27.31 | 11.86 | 324.06 | 106.99 | 2.3027
687 | 27.31 | 11.85 | 323.93 | 106.89 | 2.3046
688 | 27.31 | 11.86 | 324.02 | 107.11 | 2.3027
689 | 27.31 | 11.86 | 324.01 | 106.77 | 2.3027
690 | 27.31 | 11.86 | 323.94 | 107.17 | 2.3027
691 | 27.31 | 11.86 | 324.08 | 107.11 | 2.3027
692 | 27.31 | 11.86 | 323.94 | 107.1 | 2.3027
693 | 27.31 | 11.86 | 324.04 | 107.05 | 2.3027
694 | 27.31 | 11.86 | 324 | 107.43 | 2.3027
695 | 27.31 | 11.86 | 323.98 | 107.42 | 2.3027
696 | 27.31 | 11.86 | 324.06 | 106.97 | 2.3027
697 | 27.31 | 11.85 | 323.92 | 107.33 | 2.3046
698 | 27.31 | 11.86 | 324.06 | 107.4 | 2.3027
699 | 27.31 | 11.86 | 323.98 | 107.7 | 2.3027
700 | 27.31 | 11.86 | 323.97 | 107.75 | 2.3027
701 | 27.31 | 11.86 | 324.06 | 107.64 | 2.3027
702 | 27.31 | 11.85 | 323.93 | 107.76 | 2.3046
703 | 27.31 | 11.86 | 324.08 | 107.9 | 2.3027
704 | 27.31 | 11.86 | 323.97 | 108.13 | 2.3027
705 | 27.31 | 11.86 | 323.98 | 107.79 | 2.3027
706 | 27.31 | 11.86 | 324.05 | 108.3 | 2.3027
707 | 27.31 | 11.85 | 323.93 | 108.58 | 2.3046
708 | 27.31 | 11.86 | 324.08 | 108.61 | 2.3027
709 | 27.31 | 11.86 | 323.95 | 108.58 | 2.3027
710 | 27.31 | 11.86 | 324.02 | 108.64 | 2.3027
711 | 27.31 | 11.86 | 324.04 | 108.82 | 2.3027
712 | 27.31 | 11.85 | 323.92 | 108.74 | 2.3046
713 | 27.31 | 11.86 | 324.07 | 108.76 | 2.3027
714 | 27.31 | 11.85 | 323.93 | 108.94 | 2.3046
715 | 27.31 | 11.86 | 324.03 | 108.95 | 2.3027
716 | 27.31 | 11.86 | 324.02 | 109.5 | 2.3027
717 | 27.31 | 11.86 | 323.96 | 109.4 | 2.3027
718 | 27.31 | 11.86 | 324.08 | 109.44 | 2.3027
719 | 27.31 | 11.85 | 323.93 | 109.29 | 2.3046
720 | 27.31 | 11.86 | 324.05 | 109.48 | 2.3027
721 | 27.31 | 11.86 | 323.99 | 109.56 | 2.3027
722 | 27.31 | 11.86 | 323.96 | 109.61 | 2.3027
723 | 27.31 | 11.86 | 324.06 | 109.41 | 2.3027
724 | 27.31 | 11.85 | 323.92 | 109.7 | 2.3046
725 | 27.31 | 11.86 | 324.08 | 109.52 | 2.3027
726 | 27.31 | 11.86 | 323.99 | 109.65 | 2.3027
727 | 27.31 | 11.86 | 323.99 | 109.77 | 2.3027
728 | 27.31 | 0.01 | 0.43 | 111.81 | 2731
729 | 27.31 | 0.01 | 0.46 | 111.58 | 2731
730 | 27.31 | 0.01 | 0.5 | 111.51 | 2731
731 | 27.31 | 0.01 | 0.51 | 111.6 | 2731
732 | 27.31 | 0.01 | 0.44 | 111.42 | 2731
733 | 17.98 | 0.01 | 0.27 | 111.45 | 1798
734 | 13.99 | 0.01 | 0.26 | 111.24 | 1399
735 | 11.98 | 0.01 | 0.22 | 110.77 | 1198
736 | 11.98 | 0.01 | 0.22 | 110.54 | 1198
737 | 11.97 | 0.02 | 0.3 | 110.34 | 598.5
738 | 11.98 | 0.02 | 0.28 | 109.62 | 599
739 | 11.98 | 0.01 | 0.22 | 109.17 | 1198
740 | 11.98 | 0.02 | 0.29 | 108.39 | 599
741 | 11.98 | 0.02 | 0.28 | 107.68 | 599
742 | 11.98 | 0.01 | 0.16 | 107.11 | 1198
743 | 11.98 | 0.02 | 0.26 | 106.16 | 599
744 | 11.99 | 0.01 | 0.23 | 105.28 | 1199
745 | 11.98 | 0.02 | 0.25 | 104.26 | 599
746 | 11.97 | 0.02 | 0.29 | 103.24 | 598.5
747 | 11.97 | 0.02 | 0.25 | 102.1 | 598.5
748 | 11.98 | 0.02 | 0.28 | 101.16 | 599
749 | 11.97 | 0.01 | 0.2 | 99.95 | 1197
750 | 11.99 | 0.02 | 0.29 | 98.76 | 599.5
751 | 11.98 | 0.02 | 0.28 | 97.49 | 599
752 | 11.98 | 0.02 | 0.31 | 96.32 | 599
753 | 11.97 | 0.01 | 0.19 | 95.15 | 1197
754 | 11.99 | 0.01 | 0.22 | 93.91 | 1199
755 | 11.98 | 0.02 | 0.28 | 92.66 | 599
756 | 11.97 | 0.01 | 0.22 | 91.3 | 1197
757 | 11.98 | 0.02 | 0.29 | 90.01 | 599
758 | 11.98 | 0.01 | 0.22 | 88.8 | 1198
759 | 11.99 | 0.02 | 0.25 | 87.49 | 599.5
760 | 11.98 | 0.03 | 0.37 | 86.37 | 399.3333
761 | 11.98 | 0.01 | 0.22 | 84.91 | 1198
762 | 11.98 | 0.02 | 0.26 | 83.66 | 599
763 | 11.99 | 0.01 | 0.17 | 82.44 | 1199
764 | 11.98 | 0.01 | 0.15 | 81.1 | 1198
765 | 11.98 | 0.03 | 0.36 | 79.78 | 399.3333
766 | 11.98 | 0.03 | 0.36 | 78.5 | 399.3333
767 | 11.98 | 0.01 | 0.2 | 77.31 | 1198
768 | 11.98 | 0.02 | 0.29 | 76.08 | 599
769 | 11.98 | 0.04 | 0.51 | 74.68 | 299.5
770 | 11.98 | 0.03 | 0.41 | 73.43 | 399.3333
771 | 11.98 | 0.02 | 0.33 | 72.23 | 599
772 | 11.98 | 0.03 | 0.37 | 71.14 | 399.3333
773 | 11.98 | 0.03 | 0.42 | 70.07 | 399.3333
774 | 11.98 | 0.04 | 0.5 | 69.07 | 299.5
775 | 11.98 | 0.03 | 0.43 | 68 | 399.3333
776 | 11.98 | 0.03 | 0.43 | 67.05 | 399.3333
777 | 11.98 | 0.02 | 0.31 | 66.1 | 599
778 | 11.31 | 15.19 | 171.83 | 64.62 | 0.74457
779 | 11.31 | 15.18 | 171.77 | 63.7 | 0.74506
780 | 11.31 | 15.18 | 171.92 | 62.9 | 0.74506
781 | 11.31 | 15.18 | 171.8 | 62.24 | 0.74506
782 | 11.31 | 15.18 | 171.78 | 61.64 | 0.74506
CC mode
CR mode
CP mode
CC mode
CC mode
CR mode
CP mode
CC mode
CC mode
CR mode
CP mode
CC mode
CC mode
CR mode
CP mode
CC mode
OVT Fault
OVT Fault
OVT Fault
OVT Fault